High Impedance Active Probes

---
High-Z Active Probe
- FET amplifier in probe head
- Typically fixed HW gain
  - Some (not all) provide offset capability
- High DC input resistance (MΩ range)
- Relatively low input C (usually sub-pF)

---
High-Z Active Probe
- Most are made by scope vendors
  - Proprietary interface, not portable across makes
- Some third party ones exist
  - Tetris by PMK (1 to 4 GHz models)
  - External power supply, 50Ω BNC to scope

---
Teledyne LeCroy ZS1500
- Probe head based on PMK Tetris 1500
  - 1.5 GHz, 1MΩ || 900 fF
- ProBus control pod
  - Adds ± 12V offset capability
  - (PMK version has no offset DAC)

---
ZS1500 – S11 w/ Leaf Ground

---
ZS1500 – S21 w/ Leaf Ground

---
ZS1500 – Step Response w/ Leaf Ground

---
ZS1500: Step Response w/ 7cm Wire Ground
Worse performance due to higher L
But not actually too horrible

---
ZS1500 – S11 Across Open
PicoConnect 921
Resistive
Crossover w/ 921 at
632 MHz
Teledyne LeCroy ZS1500
Active FET
Teledyne
LeCroy PP023
R-C Divider

---
ZS1500: Time Domain Loading w/ Leaf Ground

---
High-Z Active Probe: Strengths
- Lowest DC loading of any common probe
  - Maintains high input Z much longer than R-C divider
- Relatively low capacitance
- Somewhat tolerant of poor grounding
  - Resistive probes are usually better at this, though

---
High-Z Active Probe: Weaknesses
- ESD sensitive
- Expensive (low-mid 4 digits USD)
- High attenuation (typically 10:1)
- Small but non-negligible input capacitance

---
High-Z Active Probe: When to use
- Loading-sensitive low to mid speed signals
  - The ZS1500 is my probe of choice for crystal oscillators
- Excellent general purpose embedded debug probe
  - But you can buy 4+ resistive probes for cost of one FET probe!
  - Don’t go out and buy a case of them unless you really need to