Active Differential Probes

---
Differential Probe
- Differential amplifier fed by probe tip
- Typically used for low swing, high BW signals
  - HV diff probes exist, will be covered separately

---
Differential Probe: Input Considerations
- Loading
- Common mode range
- Differential dynamic range
- Damage levels
- Overload recovery

---
Differential Probe: Grounding
- Input signal isn’t measured WRT ground, but...
- Very limited common mode range
  - ± 5V for Teledyne LeCroy DH series
  - ± 2.4V for Teledyne LeCroy D400A-AT
- Board needs to share same DC ground as probe
  - No need for a low-L RF ground path!

---
Differential Probe: Grounding Methods
- Ground from another probe
- Grounded power supply
- Ground input on amplifier
  - Only need to use one per DUT
- USB / UART / JTAG cable
- Coax shield

---
Differential Probe: Simplified Schematic

---
Loading (Teledyne LeCroy D420-A, SI tip)
Shift is almost invisible!
Good diff probes have
extremely low loading

---
Differential Probe: Tips/Accessories (D1330)
Mounting jigs, tape,
tip clips, etc.
Yes, they’re Legos!
Handheld browser
Solder-in
Pin header

---
Solder-In Tip (D1330-SI)
950Ω
100Ω
Tip is just two
resistive probes!

---
Differential Probe: Securing Tips
- Solder-in probe tips are fragile!
  - Cannot handle significant shear forces
- Secure them with tape or dedicated probe clips

---
Differential Probe: Securing Probe Body
- Amplifier is large and heavy compared to tip
- If it moves, it will damage the tip
- Use provided holders
- Taping wires down helps too

---
Overload Recovery (LeCroy D1605)
Overshoot and slow recovery
in response to sharp edge

---
Eye Pattern Comparison: 2.5 Gbps PRBS-9
LeCroy D420-A
(Active differential)
Strong overshoot
Low noise
Direct coaxial input
Best performance
High loading
Antikernel Labs
AKL-PT5
(Resistive)
Slight overshoot
Worse noise

---
Differential Probe: Strengths
- Relatively low DC loading
  - Often not as low as high-Z active probe
- Very low capacitance
- Very tolerant of poor grounding
  - Ground is only used to keep common mode in range

---
Differential Probe: Strengths
- Rejection of common mode noise
- Saves scope channels when probing a diff pair
  - Psuedo-differential input uses 2 channels + math function
- Low net attenuation (typ range /1 to /5)
  - Probe head has significant attenuation = low loading
  - Amplifier means scope sees strong signal anyway

---
Differential Probe: Weaknesses
- Very ESD sensitive
- Extremely expensive (4-5 digits USD)
- Limited range
- Overload recovery issues

---
Differential Probe: When to use
- High speed low-swing differential signals
- High speed low-swing single-ended signals (DDRx)
  - Need to make it differential somehow
  - Most common is to measure WRT ground
  - Can also measure WRT SSTL Vref etc