High Voltage Probes

---
High Voltage Probes
- Many different designs for different applications
  - HV passive
  - Fiber isolated
  - HV differential

---
High Voltage Passive Probes
- Classic R-C divider architecture
- Much higher attenuation
  - Teledyne LeCroy PPE5KV is 100:1, 5 kV max

---
High Voltage Fiber Isolated Probe
- Measuring small signal on huge DC offset
  - Teledyne LeCroy HVFO108 has 35 kV common mode range
  - Amplitude ranges from ±1V to ±40V
- Isolated front end driving optical fiber TOSA
- ROSA feeds scope input
- Often not a simple linear system
  - Fiber transceivers have poor linearity
  - HVFO108 uses FM over the fiber

---
High Voltage Differential Probe
- Active differential probe for HV applications
- Pico TA044 ($1075)
  - 70 MHz B/W
  - Switchable 100:1 / 1000:1 attenuation
  - 700 V RMS / 5 kV RMS differential range
  - 2.5 kV RMS common mode range

---
HV Probe: Strengths
Extends range of scope to several kV

---
HV Probe: Weaknesses
- Active probes often need batteries for remote head
  - Can’t be powered directly by scope
  - Some fancier ones include isolated DC-DC supplies
- Passive probes have frequency derating
  - Often can’t hit rated BW and voltage at once
- Low bandwidth (tens to low hundreds of MHz)

---
HV Probe: When to use
- Measuring high voltages
- If DC isolation is required for any other reason