Oscilloscope Probing
Andrew D. Zonenberg, Ph. D

---
Introduction

---
Today’s Class
- Lecture mixed with hands-on lab
- Stop me at any time with questions
- Bathroom / emergency procedures
- Slides are on my GitHub (CC BY-SA 4.0)
  - https://github.com/azonenberg/electronics-training/tree/master/oscilloscope-probing

---
Learning Goals
- Pros and cons of various probe designs
- How to select the best probe for a measurement
  - Not always the most expensive!
- How to get the most out of each probe
- Understand non-idealities of real world probes

---
Lecture Outline
- Introduction
- What is a probe?
- Types of probes
  - R-C divider probes
  - Resistive probes
  - Active voltage probes
  - Active differential probes
  - Power rail probes
  - Nearfield loop probes
  - Current probes
  - High voltage probes

---
About Me
- Ph.D Computer Science (RPI 2015)
- Embedded systems security by day
- High speed digital and F/OSS T&M by night
- Founder and lead developer of ngscopeclient

---
Disclaimer re Example Hardware
- Not an endorsement
  - I’m mostly a Teledyne LeCroy and Pico shop
  - Most probes we’ll discuss or use are made by them
  - This just means I have their products handy
  - Discussion of a probe doesn’t imply being particularly good / bad
- Product / model names are trademarks of their respective owners

---
Why Use One Probe Over Another?
Matrix IP1120
200 MHz
$10 on Amazon
Teledyne LeCroy D1330
13 GHz
Over $10000

---
What is a Probe?
- Both electrical and mechanical components
- Takes signal from board and puts into instrument

---
The Ideal Probe
- No influence on DUT behavior
- No noise
- No loss
- Low cost
- Unlimited frequency / voltage range
- Doesn’t exist!
  - All real probes are compromises

---
Attributes of a Probe
- What distinguishes one probe from another?
- Let’s think of as many as we can

---
Attributes of a Probe
- Bandwidth
- Attenuation
- Noise
- Flatness
- Loading
- Voltage range
- Linearity
- Cost
- Durability
- Ergonomics

---
Crash Course in S-Parameters
- We have a circuit with N ports
  - Typical probes are 2-port networks
- RF energy is applied to one port
- Some signal comes out each port
  - Outputs / reflections have amplitude and phase shift
- Model this as a NxN “scattering matrix”

---
Crash Course in S-Parameters
Notation: Sxy is path to X from Y
S21
Port 1
DUT
Port 2
S11
S22
S12

---
Crash Course in S-Parameters
- Each S-matrix element is a complex number
  - Real and imaginary
  - Or (often easier to think about) magnitude and phase angle
- Value is frequency dependent
- Nonlinear effects can create harmonics
  - S-parameters only model linear behavior
  - Keysight has X-Parameters for modeling nonlinearities.
    This is beyond the scope of today’s discussion.

---
Crash Course in S-Parameters
- Typically measured with a VNA
  - Can also simulate, etc
- Port numbering is arbitrary
  - For examples in today’s class, convention is:
    1 = DUT end
    2 = scope end
  - S11 = power reflected to DUT
  - S21 = power seen by scope

---
Direct Coaxial Connection

---
When Is A Probe Not A Probe?
When it’s just a cable!

---
External 50Ω Termination
- Lower end scopes lack native 50Ω terminations
- Can use in-line or T terminations at lower freqs

---
In-Line Terminator vs Native 50Ω Input
Matched input
No reflection
Reflection off stub between
terminator and scope frontend
Note rise time difference:
This scope is 4 GHz BW in 50Ω mode, 500 MHz in 1MΩ mode

---
Direct Coaxial Connection: Advantages
- Lowest possible noise
  - No external amplifiers
  - No attenuation so need less frontend gain
- Low cost – no expensive probe needed
- Flattest possible response
  - Only source of error is cable loss
  - Can de-embed this if cable is characterized

---
Direct Coaxial Connection: Disadvantages
- Requires 50Ω scope input
  - Inline termination works OK at lower freqs
  - Reflection issues at higher speeds
- High loading on DUT
  - Probe presents a 50Ω load
- Limited range
  - Most 50Ω scope inputs are ±5V max range
  - Many higher BW inputs are even less (±2V is common)

---
Direct Coaxial Connection: When to Use
- If your DUT already has coaxial test points
- Measuring end of unterminated 50Ω line
  - Empty DIMM or PCIe socket*
  - Card edge connector
  - Unpopulated footprint
- Ideal reference signal to compare probe against
  - It’s hard to get flatter response than a short cable