Resistive Probes

---
Resistive Probe
- The other passive probe design
  - Aka Transmission line probe, Low-Z probe, Z0 probe
- Conceptually super simple
  - Resistor and a piece of coax
- Doesn’t get as much love as it deserves

---
Resistive Probe
Linear tradeoff between insertion
loss and loading on DUT.
450Ω for 10x, 950Ω for 20x
Optional terminator
Doubles probe attenuation
but damps out reflections
if scope input is mismatched

---
Resistive Probe
- Much higher DC loading than an R-C divider
- But response can be much flatter!
  - Ideal resistive probe has C=0 and constant S21 / S11
  - Of course, parasitics ruin our fun like always...

---
Typical Resistive Probes
PicoConnect 921
20:1, AC coupled
Source terminated
$1055, 6 GHz
Pico TA061
10:1
$419, 1.5 GHz
Teledyne LeCroy
PP066
1:1 / 10:1 / 20:1
$1860, 7.5 GHz

---
Resistive Probe: S11 vs R-C Divider
R-C divider probe wins
Lower loading
PicoConnect 921
Resistive
Crossover
~100 MHz
Teledyne LeCroy
PP023
R-C divider
Resistive probe wins
-14 dB S11 @ 420 MHz
Signal on DUT is loaded down
to 20% of original amplitude!

---
Resistive Probe: S11 Crossover
Crossover is at around 100 MHz!
“500Ω” probe has lower loading
than “10MΩ” probe after here!

---
Resistive Probe: S21 Flatness Across Models
+2.7 dB peaking
PicoConnect 921
20:1, $1055
Pico TA061
10:1, $419

---
Resistive Probe: S21 Flatness Across Models
- Lower cost probe has much worse flatness
- But why?
- Let’s look at some of the effects in play

---
Chip Resistor Cross Section
NiCr film
Al2O3 brick
Solder terminals

---
But from a different perspective...
Leakage path
Dielectric
Parallel plates of conductor

---
Yep, it’s a capacitor!
Here’s a better model of a real resistor

---
Real Resistor Behavior

---
Flip-Chip Resistors

---
Resistive Probe: Ground L Sensitivity
- Much less input C than an R-C divider
  - Low end ones (Pico TA061) can be as high as 2 pF
  - Better ones (PicoConnect) are hundreds of fF
- Less C means less ringing for same L
  - You can often get away with worse grounding vs R-C divider

---
Resistive Probe: Ground L Sensitivity
Pico TA061
Spring ground
Pico TA061
Alligator ground

---
Resistive Probe: Ground L Sensitivity vs R-C
Teledyne LeCroy PP023
Alligator ground
Pico TA061
Alligator ground

---
Resistive Probe: Ground L Sensitivity vs R-C
Teledyne LeCroy PP023
Spring ground
Pico TA061
Spring ground

---
Resistive Probe: Ground L Sensitivity vs R-C
Teledyne LeCroy PP023
Spring ground
Pico TA061
Alligator ground

---
Resistive Probe: Input Stub
- Distance from probe tip to resistor is critical
  - This is an unterminated low-Z stub!
  - Reflection off resistor causes ¼ wave null in response
- Longer tip needle makes this effect worse

---
Resistive Probe: Input S11 vs Needle Length
PicoConnect 921
~6 mm tip
Pico TA061
~17 mm tip

---
Resistive Probe – Time Domain Loading (TA061)
Capacitive / stub loading from probe tip
Long needle = long reflection period
DC loading from
500Ω input resistance

---
Resistive Probe – Time Domain Loading (921)
Much smaller capacitive load!
Can still see tiny ~50ps reflection
from probe tip to resistor
DC loading comparable
to TA061

---
Resistive Probe – Input Stub
- Ideally we want no reflection at all!
  - But how can we get that?
  - Need to eliminate the mismatched stub somehow...

---
Resistive Probe – Reducing Reflections
- What if we match tip Z0 to probe resistance?
  - For example, 500Ω Z0 for a 500Ω 10:1 probe
- This eliminates mismatch at tip-to-R junction
  - All power we don’t sample is instantly reflected back to DUT

---
Resistive Probe – Reducing Reflections
- There’s just one problem…
  - The impedance of free space is ~377Ω
  - Matched tip is impossible with a 10:1 or 20:1 probe
  - Doable (maybe) with a 5:1, but mechanically tricky

---
Resistive Probe – Reducing Reflections
- Alternative: Make the stub really short
  - Move the null past the band of interest
- Lots of fun ways to do this
  - Castellated probe tips
  - Solder-in damping resistors
  - Carbon fiber tip needles
- Any of these sound familiar?
  - We’ll return to this in a later section...

---
Resistive Probe: Strengths
- Excellent price / performance ratio
  - Nothing else gives you GHz of BW for $1K!
  - Prices typically mid 3 to low 4 digits USD
- No active components – fairly ESD resistant
- Generic design – no vendor lock-in

---
Resistive Probe: Weaknesses
- Relatively high DC loading
  - Not suitable for use on lines with pullups
- May disturb DC bias on DUT
  - Can mitigate this w/ coaxial DC block
  - Some probes include AC coupling cap
- High attenuation (typically 10:1 or 20:1)
  - Can reduce attenuation at the cost of worse DC loading

---
Resistive Probe: When to Use
- Fast digital signals with push-pull drivers
  - Excellent general purpose embedded debug probe
  - Easily usable out to several Gbps
- Low impedance analog
  - Great for 50Ω RF