High Impedance Active Probes

---
High-Z Active Probe
- FET amplifier in probe head
- Typically fixed HW gain
  - Some (not all) provide offset capability
- High DC input resistance (MΩ range)
- Tiny input C (usually sub pF) vs passive R-C

---
High-Z Active Probe
- Most are made by scope vendors
  - Proprietary interface, not portable across makes
- Some third party ones exist
  - Tetris by PMK (1 to 4 GHz models)
  - External power supply, 50Ω BNC to scope

---
Why put an amplifier in the probe?
- Typical scope input: 10MΩ || 17 pF
- Compare to a modern JFET buffer IC
  - Older probes usually used discretes
  - Same general architecture

---
Not so fast...
- 50GΩ || 2.4 pF sounds great at first
  - But that’s still more capacitance than we’d like
  - And 50GΩ is overkill for almost all real world circuits

---
Can we hide some of this capacitance?
- Capacitors in series divide
  - But we still need a DC path :(
  - What about a parallel R?
- R || C between input and amplifier
- ~1MΩ termination R across amplifier input

---
That looks familiar...
- Yep, it’s our old friend the R-C divider probe!
- But matched to 2.4 pF amplifier, not 17 pF scope
  - This means we need much less series C
  - AKL-AV1: 5MΩ || 350 fF equivalent

---
Other design considerations
- High BW JFET amplifiers often come w/ large offset
- May need to combine two
  - High BW amplifier for AC path
  - Precision amplifier for DC path
- Distance from probe tip to amplifier
  - Small damping resistor (100Ω) to reduce reflection

---
Putting it all together: AKL-AV1 schematic 1/2
AC path amplifier
DC path amplifier
R-C divider

---
Putting it all together: AKL-AV1 schematic 2/2

---
Teledyne LeCroy ZS1500
- Probe head based on PMK Tetris 1500
  - 1.5 GHz, 1MΩ || 900 fF
- ProBus control pod
  - Adds ± 12V offset capability
  - (PMK version has no offset DAC)

---
FIXME REDO
ZS1500 – S11 w/ Leaf Ground

---
ZS1500 – S21 w/ Leaf Ground

---
ZS1500 – Step Response w/ Leaf Ground

---
ZS1500: Step Response w/ 7cm Wire Ground
Worse performance due to higher L
But not actually too horrible

---
FIXME REDO
ZS1500 – S11 Across Open
PicoConnect 921
Resistive
Crossover w/ 921 at
632 MHz
Teledyne LeCroy ZS1500
Active FET
Teledyne
LeCroy PP023
R-C Divider

---
ZS1500: Time Domain Loading w/ Leaf Ground

---
High-Z Active Probe: Strengths
- Lowest DC loading of any common probe
  - Maintains high input Z much longer than R-C divider
- Relatively low capacitance
- Somewhat tolerant of poor grounding
  - Resistive probes are usually better at this, though

---
High-Z Active Probe: Weaknesses
- ESD sensitive
- Expensive (low-mid 4 digits USD)
- High attenuation (typically 10:1)
- Small but non-negligible input capacitance

---
High-Z Active Probe: When to use
- Loading-sensitive low to mid speed signals
  - The ZS1500 is my probe of choice for crystal oscillators
- Excellent general purpose embedded debug probe
  - But you can buy 4+ resistive probes for cost of one FET probe!
  - Don’t go out and buy a case of them unless you really need to

---
Questions?